RDVLF-Z80Y  Very Low Frequency HV Test Set

The withstand voltage test is divided into two categories: AC and DC. The AC withstand voltage includes technologies such as power frequency, variable frequency, and 0.1Hz very-low frequency. As the latest testing solution recommended by the International Electrotechnical Commission (IEC), the 0.1Hz very-low frequency technology has significant technical advantages.

RDVLF-Z80Y features auto-variable frequency (0.1–0.01Hz) true sine wave output, with a color touchscreen that displays voltage and current waveforms in real time. It incorporates safe optical isolation, closed-loop negative feedback, and overvoltage/overcurrent protection. Direct sampling on the HV side ensures accurate data. Built-in printer and protective resistors are included, with support for result storage and printing. Portable and rugged, it is ideal for on-site AC withstand voltage testing.

- Product features

- AC Withstand voltage testing, Voltage & current curve display.

- Safe Optical isolation technology.

- Easiest to use Fully automatic microcontroller.

- H/L voltage closed loop negative feedback control circuit, output no capacity rise effect.

- Built-in printer.

- Built-in protective resistors.

- Color Touch LCD (154*86mm) operation, True sine wave output.

- Real-time display of actual output waveform.

- Ergonomic, menu guided, large color touch screen

- Storage of test results for later retrieval and print.

- Over voltage and current protected.

- Auto variable frequency 0.1Hz, 0.05Hz, 0.02 Hz, 0.01 Hz.

- 60 mins continuous working;

- Data of current, voltage and waveform directly sampled at HV side.

- Portable rugged and reliable trunk design (box material is special mixed resin, made in USA).

- RS232/USB for buyer’s option (Default configuration of USB interface).

- DC withstand voltage test and DC leakage current test function (optional).

- Product specifications and technical parameters

- Power: 220V ± 10% AC (110V can be customized), 50/60Hz, single phase;

- AC output-Voltage (Peak): 80 KV

- Output frequency: 0.1Hz, 0.05Hz, 0.02 Hz, 0.01 Hz (auto variable frequency);

- Load: 0.5μF @ 0.1 Hz; 1.0μF @ 0.05 Hz; 2.0μF @ 0.02 Hz; 4.0μF @ 0.01 Hz

- Stability of frequency: Fluctuation < 0.5%;

- AC voltage measurement accuracy: ±3%;

- AC voltage measurement resolution: 0.1KV;

- AC current measurement accuracy: ±3%;

- AC current measurement 0-20mA, resolution: 0.1mA;

- Positive and negative voltage peak error: ≤3%;

- Voltage waveform distortion: ≤3%;

- Construction: 1*40KV integrated device + 1*40KV external booster

- DC leakage current measurement (optional)

DC output rated voltage: +80kV           DC voltage accuracy: 3%

DC leakage current range: 0-2000μA       DC leakage current resolution: 1μA

DC leakage current accuracy: 3%

- Ambient conditions of use: temperature Indoor and outdoor : -10ºC ～ + 50ºC.

- Humidity: ≤85%