STEADY-STATE SOLAR SIMULATOR

Technical Specification

Model: LC-LED-AAA-2400

Chapter 1: Product Overview

1.1 Test Items and Conditions

Test Standard: IEC 61215-2:2021 MQT-2

Test Conditions:

Temperature: 25°C

Irradiance: 800-1200 W/m²

1.2 Applicable Standards

IEC 60904-9:2020 - "Photovoltaic devices – Part 9: Classification of solar simulator characteristics"

Chapter 2: Technical Specifications

| No. | Parameter | Specification |
| --- | --- | --- |
| 1 | Solar Simulator Classification | Class AAA |
| 2 | Spectral Match (AM1.5G) | Class A+ |
| 3 | Irradiance Spatial Non-uniformity | ±2% (Class A) |
| 4 | Irradiance Temporal Instability (Long-term) | ±2% (Class A) |
| 5 | Irradiance Intensity | 1000 W/m² |
| 6 | Effective Illumination Area | 2400mm × 1200mm |
| 7 | Light Source Type | Metal Halide Lamp (MHL) |
| 8 | Optical Path Length | >10m |
| 9 | Light Emission Direction | Side lighting |
| 10 | Light Source Cooling | Forced air cooling |
| 11 | Module Temperature Control | Independent temperature control in test zone |
| 12 | Total Power | Approx. 100 kW |

System Design

The system employs a side lighting design with independent air-conditioned temperature control zones on both sides.

Full-area simultaneous illumination designed for large-format PV module testing.

Chapter 3: Key Technical Implementation

3.1 Light Source (MHL)

Spectral Performance:

Wavelength range: 300nm - 1200nm

Meets IEC 60904-9:2020 Class A requirements

6-segment spectral match: Class A

Spectral Distribution:

Spectral Match Table (AM1.5G Standard):

| Wavelength Range (nm) | AM1.5G Reference (%) | Achieved Match | Class |
| --- | --- | --- | --- |
| 300-470 | 16.61 | Within tolerance | A |
| 470-561 | 16.74 | Within tolerance | A |
| 561-657 | 16.67 | Within tolerance | A |
| 657-772 | 16.63 | Within tolerance | A |
| 772-919 | 16.66 | Within tolerance | A |
| 919-1200 | 16.69 | Within tolerance | A |

3.2 Collimation and Irradiance Uniformity

Technology:

Concentrated light projection with overlapping layers

Achieves Class A (±2.0%) irradiance non-uniformity

Uniformity Test Results:

Key Advantages:

Overlapping Projection Method: Avoids non-uniformity caused by light spot overlap in tiled array configurations; Eliminates light energy loss from edge light spillage

Optimized Light Spot Design: Each light source projects a 2:1 rectangular spot; Projection area: 2400mm × 1200mm; Overlapping projection achieves cumulative irradiance ≥1000 W/m²

Long Optical Path: Effectively eliminates thermal radiation from light source to samples; Minimizes edge light loss

Irradiance Spatial Non-uniformity: Class A

Variable Focus Projection:

Adjustable projection focus for optimal uniformity control

Assembly Workshop:

3.3 Incidence Angle Compliance

Standard Requirement:

For solar simulators used in power testing, the incidence angle shall not exceed 15°, as specified in JIS C 8933-1995(2011) "Solar simulators for amorphous solar cells and modules"

Our Design:

Complies with incidence angle requirements

Not limited by illumination area size

Capable of optimizing key parameters: spectral match, non-uniformity, spectral distribution uniformity, spectral coverage, spectral deviation, spectral mismatch factor

Chapter 4: System Specifications Summary

4.1 Performance Parameters

| Parameter | Value |
| --- | --- |
| Effective Illumination Area | 2400mm × 1200mm |
| Irradiance Intensity | 1000 W/m² (1 sun) |
| Spectral Match (AM1.5G) | Class A+ |
| Spatial Non-uniformity | Class A (≤2%) |
| Temporal Instability | Class A (≤2%) |
| Wavelength Range | 300nm - 1200nm |
| Optical Path | >10m |
| Light Source Lifetime | >1,000 hours |
| Total Power | ~100 kW |
| Cooling Method | Forced air cooling |

4.2 Standards Compliance

IEC 60904-9:2020 - ✅ Class AAA

IEC 61215-2:2021 MQT-2 - ✅ Suitable for testing

JIS C 8933-1995(2011) - ✅ Incidence angle compliant

4.3 Included Components

MHL steady-state solar simulator main unit (2400×1200mm)

Air cooling system for light source

Temperature control system for test zone

Control panel and remote operation interface

Installation and operation manual

Calibration certificate (IEC 60904-9)

4.4 Optional Accessories (Not Included)

IV test module

Source meter

Data acquisition system

Sample stage

Environmental chamber integration

Chapter 5: Warranty and Service

5.1 Warranty

Standard warranty: 12 months from delivery

Covers manufacturing defects and component failures

5.2 Technical Support

Remote technical support during warranty period

On-site service available (travel costs apply)

5.3 Delivery and Installation

Delivery time: To be confirmed based on production schedule

Installation and commissioning: Included

Training: Included

────────────────────────────────────────────────────────────

Document Version: 1.0

Date: May 2026