Technical specifications

| Product Name: | Multi-channel photovoltaic module steady-state test system |
| --- | --- |
| Product Model: | MCT10-300-20T |

Lecheng Intelligent Technology (Suzhou) Co., Ltd

December 2024

directory

1. Equipment overview	3

2. Test object	3

3. Specification summary	3

1. Technical parameters	3

2. Equipment structure	3

3. Main components of the equipment	4

4. Main component functions	4

1. Steady-state light source	4

2. Multi-channel electrical test cells	6

3. Multi-channel temperature control and monitoring module	6

4. Software testing systems	7

5. other	8

5. Factory business needs	8

6. Annex related	8

1. Documentation	8

2. Amenities	8

3. Consumables	9

7. Service description	9

1. Installation and commissioning	9

2. After-sales training	9

3. Equipment maintenance	9

- Equipment Overview:

MCT series equipment is a standard solar simulator that irradiates multiple perovskite photovoltaic modules at the same time, controls the temperature of each photovoltaic module separately through multiple channels, and collects the I-V performance parameters of each module and the MPPT maximum power point tracking data at the same time, so as to achieve high-throughput testing of the temperature stability and light operation stability of photovoltaic modules, providing convenience for the development of perovskite materials and process reliability screening.

The equipment is efficiently integrated, integrating LED steady-state simulator light source, multi-channel test base, switching platform, multi-channel temperature control and monitoring module, multi-channel test source table and load, data acquisition display automatic storage system, etc., which can achieve high throughput, independent non-interference of each module, long-term automatic test of high temperature and light operation performance degradation data of batteries and modules.

- Test objects

The test objects are thin-film solar cells and modules, and the substrate is glass or flexible materials, and the specific specifications are as follows.

| project | numeric value |
| --- | --- |
| Material | Glass substrates or flexible substrates thin-film solar cells and modules |
| thickness | 1mm—8.0mm |
| Test sample size | 50*50mm, 158*158mm, 220*220mm, 300*300mm package components, positive and negative lead wires |

- Specification overview

- Technical parameters

| Model | MCT10-300-20T |
| --- | --- |
| Adapts to product size | 50*50mm、158*158mm、220*220mm、300*300mm |
| Device size | 1200×800×1800mm |

- Equipment structure

The equipment is composed of light source module, test platform, multi-channel test unit, temperature control module and control software. The light source is a down-lit structure, and the height of the test platform is controlled by the lifting module, and the test platform can be moved left and right to switch between multi-channel test and single large component test mode.

- Main components of the equipment

| serial number | fittings | quantity | remark |
| --- | --- | --- | --- |
| 1 | Class 3A LED steady-state solar simulator | 1 set | 100*100、220*220、300*300mm |
| 2 | Multi-channel electrical test cells | 1 set | 1-channel/20-channel switching, 4-20 channels can be customized |
| 3 | Multi-channel temperature control and monitoring module | 1 set | 1-channel/20-channel switching, 4-20 channels can be customized |
| 4 | Multi-channel I-V, MPPT data acquisition, automatic monitoring data storage software system | 1 set | The interface and functions can be customized according to your needs |

- Main component functions

- Steady-state light source

The light source adopts a 3A LED steady-state solar simulator.

| Model | LED-3A-100 | LED-3A-220 | LED-3A-300 |
| --- | --- | --- | --- |
| Light source category | LED steady-state light source |  |  |
| Light source life | 10000h |  |  |
| Effective area | 100*100mm | 220*220mm | 300*300mm |
| Spectral wavelength range | 350-1100nm、300-1100nm、350-1200nm、300-1200nm可选 |  |  |
| Power adjustment range | 100-1100 W/m2 adjustable |  |  |
| Spacing adjustable range | 10-40cm, accuracy 0.1cm |  |  |
| Spectral match | A/A+（AM1.5G） |  |  |
| Time instability | A≤±2% |  |  |
| Spot uniformity | A≤2% |  |  |
| Supply voltage | Power Supply: AC 220V/50HZ |  |  |

- Multi-channel electrical test cells

The multi-channel electrical test unit can test 20 channels at the same time, the component size is basically 50*50mm, and other size specifications can be customized. At the same time, the test platform can be directly switched to a single channel of large modules, which is compatible with various types of crystalline silicon modules 156/182/210mm and corresponding encapsulated crystalline silicon/perovskite stacked modules. The electrical parameters are as follows:

| Model | Multi-channel (perovskite or stacked) DC-PS/PT | Single channel (crystalline silicon or stackup) SC-SS/PT | Single-channel perovskites SC-PS |
| --- | --- | --- | --- |
| Multi-channel test base | 4/8/20 channels within 50mm | 156/182/210mm | 300mm |
| Maximum voltage range | 10V; 18V | 80V | 100V |
| Minimum voltage range | 1V | 10V | 300mV |
| Maximum current range | 0.5A; 1A | 20A | 1A |
| Minimum current range | 5uA | 1mA | 100nA |
| Measurement accuracy | 0.1% | 0.1mV/0.1mA ±(0.025%+0.025%FS) | 0.1% |
| Power accuracy | 10W/CH | 250W/CH | 30W/CH |

- Multi-channel temperature control and monitoring module

The base has more than 20 channels of independent temperature control function (25-100 degrees), and multi-channel monitoring of the temperature function on the back of the module, and the real-time feedback of the chip thermocouple ensures that the temperature on the back of the module remains stable, and the handle design can be taken independently at any time without interfering with each other.

| Model | Multi-channel (perovskite or stacked) DC-PS/PT | Single channel (crystalline silicon or stackup) SC-SS/PT | Single-channel perovskites SC-PS |
| --- | --- | --- | --- |
| Multi-channel test base | 4/8/20 channels within 50mm | 156/182/210mm | 300mm |
| Temperature control range | Each channel is independently controlled at 25-100°C (±2°C) with temperature feedback |  |  |
| Temperature monitoring accuracy | 0.1°C |  |  |
| Temperature curve display | It can be displayed and saved in real time |  |  |

- Software testing systems

The equipment software is independently developed by Lecheng Intelligence, based on strong equipment integration capabilities, and the company can carry out various customized designs according to customer needs when the hardware configuration meets the conditions.

The main functions of the software system include multi-channel I-V testing, which can realize automatic continuous scanning and interval time testing, and each channel is controlled independently; Multi-channel MPPT testing, multi-channel temperature curve monitoring, timing test and timing save function, and the small device unit can automatically inspect and test a single pixcel point.

| Model | Multi-channel (perovskite or stacked) DC-PS/PT | Single channel (crystalline silicon or stackup) SC-SS/PT | Single-channel perovskites SC-PS |
| --- | --- | --- | --- |
| I-V test mode | forward and reverse scanning test; 9 points are proposed to be legal; Dynamic IV |  |  |
| MPPT tracking algorithm | disturbance observation method; conductance increment method; Fixed voltage method |  |  |
| Monitor data types | Voc、Isc、FF、PCE、Pmax、Imax、Vmax、Rs、Rsh |  |  |
| Data storage | automatic naming and scheduled saving; Each channel is saved independently or unified |  |  |
| Temperature curve display | It can be displayed and saved in real time |  |  |
| Multi-channel control | Each channel has independent start and stop, independent control of parameters, and independent control of temperature |  |  |

- other

| project | illustrate |
| --- | --- |
| Cooling | Air cooling is used to control the ambient temperature |
| Humidity control | A humidity control module can be added to control the humidity in the test area |
| dustproof | Seamless contact is used to isolate dust influences |

- Factory needs

| power supply | 220 V AC、50 Hz、10 A |
| --- | --- |
| Device size | 1200×800×1800mm |
| Ambient temperature | 22 – 27℃ |
| Ambient humidity | 20% — 80% |

- Annex related

- Documentation

When the equipment is delivered, the list of accompanying materials is as follows:

| serial number | Name of the data | Quantity (servings) |
| --- | --- | --- |
| 1 | Cargo packing list | 1 |
| 2 | Equipment instruction manual | 1 |

- Amenities

For the key components of the equipment, Lecheng has built a spare parts warehouse, which can be replaced if there is a major failure on site, the list is as follows:

| name | Model specifications |
| --- | --- |
| Multi-channel temperature control module |  |
| Multi-channel test module |  |
| Substrate fixtures |  |

- Consumables

Commonly used consumables for equipment are as follows:

| name | Replacement cycle |
| --- | --- |
| terminal blocks | 1 year |
| Heating plates | 1 year |

- Service description:

- Installation and commissioning

According to the contract, the company safely transports the equipment to the installation location designated by the user within the specified time, and sends technical service engineers to install it on site. In the case that the user has basic installation and debugging spare parts, the technical service engineer will complete the installation and debugging of the equipment for the user to use within 3 working days to ensure that the installation and commissioning site environment is neat, clean and orderly.

- After-sales training

After the equipment is debugged, the company conducts no less than 1 day of technical training for user operators, and the main training content is as follows:

- Multi-channel test software usage training;

- Multi-channel test platform operating specifications;

- training on and off operating procedures;

- the significance of panel and software control parameters, training on the range of parameter selection;

- basic cleaning treatment and maintenance of the machine;

- Common hardware fault handling and operation issues that should be paid attention to.

- Equipment maintenance

From the date of acceptance of the equipment, Lecheng equipment provides a free warranty period of twelve months for the whole machine, and provides lifelong maintenance services.

During the warranty period, after-sales personnel will be arranged to arrive at the customer's site every three months to inspect the equipment, provide basic maintenance of the equipment and give relevant usage opinions.

If the equipment fails, the after-sales response time during the warranty period is within 24 hours, and the user engineering personnel will be given remote guidance as soon as possible through telephone, video, etc.; For emergency faults that cannot be solved on-site by user engineering personnel, the company's maintenance personnel will arrive at the site to solve them within 72 hours; The maintenance of equipment outside the warranty period shall be agreed separately by both parties.