SPECIALIZED IN MEASURING

LIQUID COLORIMETRY

Accurately captures the subtle color differences in liquids.

[image: image1.png]



Note: This model does not have a display screen and requires connection to computer software for measurement control.

1、Product Introduction

The Liquid Colorimeter TS4003 is a professional color analysis instrument.Its application software supports calculations of various color spaces and color difference formulas.It can also measure transmittance and a variety of colorimetric indicators,meeting the diverse needs of different industries.The application areas cover glass processing,plastic processing,film,display processing,packaging industry,and liquid pharmaceutical analysis,among others.
2、Product Features

1. Rich Colorimetric Indicators

In addition to the measurement of common color indicators, it also provides a variety of colorimetric indicators such as transmittance, ASTM D1500 Petroleum Color, Platinum-Cobalt Index, Gardner Index, Saybolt (Saybolt Index), Pharmacopoeia (Chinese Pharmacopoeia, European Pharmacopoeia, United States Pharmacopoeia), EBC (Beer Color), and ICUMSA Color (Sugar Color), meeting the diverse needs of various industries. 

[image: image2.jpg]Pdew

Color Index Transmittance ASTM D1500 Pharmacopoeia EBC Beer
Measurement Petroleum Color
Color Scale





Intelligent Zero Calibration

Utilizing intelligent zero calibration technology, it adapts to various measurement environments, maintaining the instrument's optimal performance. This ensures the accuracy of measurement data and high consistency of measurement results.  

[image: image3.jpg]Repeatability No intelligent zero calibration technology -

- large data fluctuation

0.1

Intelligent zero calibration technology - data stability
0.01

1 2 4 24 120 1200 3600 Time (h)




Outstanding Performance

 Superior Repeatability and Inter-instrument Difference: Spectrophotometric transmittance repeatability ≤ 0.05%, color value repeatability ≤ 0.015 ΔE, and inter-instrument difference ≤ 0.25 ΔE among multiple devices, ensuring consistent quality control.

Fast Detection Efficiency: A single measurement takes only 0.8 seconds. Combined with the automatic calibration function, it can continuously test over 300 samples per hour, significantly improving measurement efficiency.

[image: image4.jpg]Single measurement takes only
can be continuously

tested per hour.

Spectrophotometric transmit-
tance
Color value repeatability





4.Diverse Light Source Simulation

The instrument is equipped with a series of standard light sources recognized by the International Commission on Illumination (CIE). By adjusting the test light sources through the software, it ensures precise color measurement of your products under different lighting conditions.
[image: image5.jpg]



5.Full-Spectrum Precision Analysis

Adopting the D/0 optical geometry (diffuse illumination, 0° viewing) and a full-spectrum LED light source (360-780nm), in conjunction with a dual-array silicon photodiode sensor, it enables simultaneous measurement of transmittance, colorimetry, color difference, and spectral curves. This achieves high-precision measurement.
[image: image6.jpg]d:0°

Sensor

—

Test the optical Splitter path system
path system v

sample

Integrating sphere

&

illuminant

‘// Light barrier

D/0 optical geometry

121 Relative spectrum

“ Missing Spectral

Dual-array silicon photodiode sensor

127 Relative spectrum

h Spectral equalization

30 40 s0 w0 700 a0
Wavelength(nm)

3 40 s o 70 &0
Wavelength(nm)

Full-spectrum LED light source




6. Convenient Accessory Expansion

The instrument comes standard with 10mm and 50mm cuvettes and is compatible with 30mm and 100mm specifications. It can accommodate different sample volumes to meet personalized needs. [image: image7.jpg]Adaptable to different sample volumes

10mm 30mm 50mm 100mm




PC Management Software

It provides a powerful PC-based management and control software, enabling stable and reliable industrial-grade instrument control. Suitable for quality management in various industries, it generates test reports from users' measurement results, facilitating customized and efficient management for customers.

[image: image8.jpg]PC quality management software





3、Product parameter

	Model number
	TS4003((Screenless version,with functions switched through computer software)

	Lighting mode
	Transmittance: D/0 (Diffuse Illumination, 0° Direction Detection);

Approximates CIE15 Standard 0/0 (Parallel Light Illumination/0° Detection);

Complies with Standards: ASTM D1003/1044, CIE No.15, GB/T 3978, GB 2893, GB/T 18833, ASTM E308, DIN5033 Teil7

	Characteristic
	Applications:Widely used in industries such as glass processing, plastic processing, films, display manufacturing, packaging, and liquid pharmaceutical analysis for colorimetric analysis and transmittance testing.

	Standalone Operation
	/

	Integrating Sphere Size
	Φ40mm

	Lighting source
	Combined full spectrum LED light source

	Spectroscopic method
	Planar grating spectroscopy

	Sensor
	Silicon photodiode array  (double row 40 groups)

	Wavelength Range
	360-780nm

	Wavelength Interval
	10nm

	Semiband Width
	5nm

	Transmittance 

measurement range
	0~200%

	Color Space
	CIE LAB,XYZ,Yxy,LCh,CIE LUV,Musell,s-RGB,HunterLab,βxy,DIN Lab99 

	Color Difference Formula
	ΔE*ab,ΔE*uv,ΔE*94,ΔE*cmc(2:1),ΔE*cmc(1:1),ΔE*00, DINΔE99,ΔE(Hunter),ΔE*CH

	Other Colorimetric Index
	Transmittance (T), Whiteness Index (WI) (ASTM E313, CIE/ISO, AATCC, Hunter, Taube, Berger, Stensby), Yellowness Index (YI) (ASTM D1925, ASTM E313), Tint Index (TI) (ASTM E313, CIE/ISO), Metamerism Index (Mt), Colorfastness to staining, Colorfastness to light, Absorbance, APHA/Hazen/Pt-Co (Platinum-Cobalt Index), Gardner Index, Iron-Cobalt Index, Saybolt Index，ASTM D1500 Color Scale, Pharmacopoeia (Chinese Pharmacopoeia, European Pharmacopoeia, US Pharmacopoeia), EBC (Beer Color), ICUMSA Color (Sugar Color)

	Observer Angle
	2°/10°

	Illuminant
	D65,A,C,D50,D55,D75,F1,F2,F3,F4，F5，F6,F7,F8,F9，F10,F11,F12,CWF,DLF,TL83,TL84,TPL5,U30，B,U35,NBF,
ID50,ID65,LED-B1,LED-B2,LED-B3,LED-B4,LED-B5,LED-BH1,LED-RGB1,LED-V1,LED-V2,LED-C2,LED-C3,LED-C5

	Displayed Data（Display on PC software)
	Spectral graph/data, Sample colorimetric values, Color difference values/graphs, Chromaticity diagram, Color simulation,Pass/Fail results, Color bias

	Measuring Time
	Rapid mode: approximately 0.8 seconds

	Repeatability
	Spectral Transmittance: Standard deviation within 0.05% (within 0.04% for 400-700nm)• Colorimetric Values: ΔE*ab within 0.015 (after preheating and calibration, measured at 5-second intervals)

	Inter-instrument Error
	Inter-instrument difference: ΔE*ab within 0.25

	Dimension
	Length × Width × Height = 370 × 300 × 200 mm

	Weight
	Approximately 9.6 kg

	Power
	DC 24V/3A power adapter

	Illuminant Life Span
	Over 3 million measurements within 5 years

	Display
	/

	Data Port
	USB

	Data Storage
	Large-capacity storage via a host computer

	Special function
	Intelligent zero calibration

	Language
	/

	Operating Environment
	0~40℃（32~104°F）

	Storage Environment
	 -20~50℃（-4~122°F）

	Standard Accessory
	Power adapter, Quality management software (downloadable from the official website), Data cable, Transmittance black plate, 10mm cuvette, 50mm cuvette, 50mm cuvette holder

	Optional Accessory
	Microhole (4mm) transmittance test fixture assembly, 30mm cuvette, 100mm cuvette

	Note:
	Subject to change without notice.


